## Slide 1: SpikeGLX
Highlights, hints and tips
UCL Course 10/14/2024
Notes: This is a quick overview of key software highlights. On a lot of slides I’ll just point you to other references.

## Slide 2: Resources (all free, open source)
Download page, it’s all here: https://billkarsh.github.io/SpikeGLX/
Get the latest version: bug fixes, richest metadata
Quick-ref: tour all the main features
Videos
Manuals and help documents
Metadata Guides, Latest items:
  GeomMaps
  Tip Length
  Anatomy
  Probe error flags
  Mux Table (helps 3rd-party software do multiplexing time-shift corrections)
Metadata complete: Everything you need to know about THIS probe is in there… 3rd-party processing tools don’t need tables of auxiliary data.
Notes: Start at the download site, it has everything you need:
Quick-ref
Help docs
Videos
There’s a lot of metadata stored with each data acquisition session:
All the hardware parts you used
All the settings and parameters
So you can recreate what you did and track down any parts that might have problems

## Slide 3
Topics on this page
It scrolls!
Click Help to see the QuickRef
Content Tabs
Change history here
https://billkarsh.github.io/SpikeGLX/
Notes: Here’s how the download site looks.
Major topics across the top
Table of contents on the left
Scroll through document on the right.
Where to get stuff
How to install and use it

## Slide 4: Probe Geometry (Metadata ~snsGeomMap)
Z=0
tip length
X=0
X=0
X=0
X=0
width
shank spacing
Z
X
Replaces ~snsShankMap.
Gives locations of all selected sites on the probe surface in microns.
Uses coordinate system shown here.
This is described in ‘Metadata_30’ guide
Notes: Imec probes have many more sites than recording channels, so you usually record from some subset of the sites.
The GeomMap is in your recording metadata. It gives the physical locations in microns of the sites you actually used.
Look on the download site document: Metadata_30 for details.

## Slide 5: How to Run
New Acquisition (Ctrl+N)
Configure dialog ->
Data flow: Follow the tabs:
  Which hardware
  Hardware parameters
  Synchronizing the hardware
  How/when files are written
  Where to save data, what to call it
Context help at each step
Notes: To configure a new data acquisition run you use the Configure dialog.
Select File::New Run
In the Configure dialog, go through the tabs at the top in order.
There’s help for each tab.
When you’ve set everything, click Run at the bottom.
The dialog remembers everything you did last time, so really, you only need to review settings…and maybe change one or two.

## Slide 6
Where I plugged my probes,
Which {slot port dock}. NEW: Slots can be PXI, OneBox, Sim.
All NI devices are here
Help on current tab
? in title bar = more help
Notes: This is the first tab, called Devices.
Here you check boxes to select which hardware you want to run, and to indicate where you plugged in your probes.
Then you click Detect which talks to all the devices you selected to see if they are attached and working. All the part numbers and serial numbers are read from the devices at this time so you can see what’s plugged in.
Also note the two buttons here:
Configure Slots
Sim Probes
Config slots is needed for the new OneBox
Sim Probes lets you run any existing recording back through SpikeGLX to simulate a real run.

## Slide 7: OneBox ! ?
8 x 8 X 1 inches
USB 3.0
2 probe ports
12 Aux analog channels in
12 analog out
~ 3600 EUR
Order: Neuropixels.org
Quickstart Guide:
SpikeGLX/_Help
Online
Notes: These are the main specs for OneBox.
We’ve written a nice getting started doc that comes with SpikeGLX.
Note you can run more than one box at a time
You can run a combo of PXI and OneBox
SpikeGLX will soon add support for analog output.
We will have a stimulus waveform designer
We will add triggered play/repeat cycles for waveforms.

## Slide 8
Assign slot to each OneBox!
Add…
Browse for run file
Assign { slot port dock }, double-click the cells
Click Enable
Devices Tab
Notes: The two buttons on the DEVICES tab.
Use Configure slots to assign a slot to each OneBox.
For each slot, show either one or two docks per headstage.
The other dialog is for Simulated Probes.
Select any run file
Attach it to a simulated slot address, or, a real probe address
The data for that (slot,port,dock) comes from that file.

## Slide 9: IM Setup Tab
Lower latency remote fetching NEW!!
Survey mode
  Check box, secs/bank
  FileViewer to see survey
Filtered buffers NEW!!
  Check box, bandpass
  Better SNR:
    Audio
    ShankViewer
    SpikeViewer
IMRO editor
  Double-click cell
See, context sensitive
Notes: Low latency runs the acquisition unit harder to reduce latency, but may be too much CPU activity if using lots of probes.
Click the box to run a SURVEY, set how many seconds per bank (solid set of 384 channels), and can set maximum bank if probe not inserted all the way. That’s it.
The filtered queues are used by audio and the ShankViewers to get better signal to background. You can also fetch these data from the remote APIs.

## Slide 10
Set rows/box
Set width {full, half} shank
Click to place centers of boxes
Resize; shift-click to remove box
When sum = rqd, save as a file
But this is much cooler when you can see the activity on the shank…
And you can:
Online in ShankViewer
Offline in FileViewer
ShankViewer Edit tab!!
IMRO Editor
Green boxes select sites for readout
Notes: Much more flexibility with selection.
- Full width or Half width boxes.
- Boxes can be resized.
- Boxes can be different sizes.
Save channel subset can now be set here.

## Slide 11
Right-click to see
more commands
Notes: A bit hidden, SHANKVIEWER.
Has two tabs.
Edit the site selection IMRO table.
View activity
Also, right-click on trace or site to get other features…
NEXT SLIDE

## Slide 12
Right-click menu
Notes: Things you can do from right-click.
Audio
SpikeView
Mask channel
Change save subset
Get ColorTTL feature.

## Slide 13
ColorTTL
Track up to 4 TTL channels
Each a different color
Notes: See annotation on slide

## Slide 14
Note:
Anatomy simulated for illustration purposes
Anatomy stuff is on the ShankViewer ‘View’ tab.
Notes: To see the ANATOMY data coming live from Pinpoint or Trajectory Explorer…
Start run
Open ShankViewer
View Tab displays the region legend/color table
Apply color to shank or traces

## Slide 15
Link
Ctrl+click = zoom
Shift+click = meas. span
Notes: A key FileViewer tool is LINKING.
You can open a set of files and time lock them. Scroll through them together to see how events are aligned in time.
Also we point out that you can:
Control-click to select a time range and zoom in on it.
You can also change the time scale with ctrl +/-
Shift-click to select and measure a time span. Read out the span in the bottom status region of the window.
CLICK TO PLAY MOVIE

## Slide 16
Open survey
Integrate activity
Select hotspots
Notes: Here we review the results of a probe SURVEY.
Open the recorded file with FileViewer.
You can scroll manually to see how stimulus events align with activity; LINKING.
Use ‘Sample Whole Survey’ + ‘Update’ to integrate all activity.
Now you can create imto tables that lasso the activity,
Or intersection of activity and anatomy,
Or increase/decrease in activity around stimulus markers in a TTL channel, which is coming soon.
CLICK TO PLAY MOVIE

## Slide 17
Notes: Also within the FileViewer is the Export tool
Downsize a large data file
Any subset of the channels
Any subset of the time range
This makes a new smaller run you can:
Test and debug issues
Share easily with others

## Slide 18: Other SpikeGLX Features
Windows::More Traces (ctrl+T) : opens a second Graphs window with two more streams and ShankViewers.
Windows::Run Metrics (ctrl+M) : Monitors performance of SpikeGLX.
Graphs Window lower-right menu : select layout of trace viewers.

## Slide 19: Near Term SpikeGLX Features
Click-to-listen shortcut
Waveform designer + triggered playback (OneBox, NI)
PSTH for opto-tagging
Support ONYX Commutator + IMU headstage

## Slide 20: Scripting (‘SDKs’)
C++
C
C#
Python
MATLAB
Control everything
Get real-time data, extend SpikeGLX feature set
Mean 4.8
Mean 3.6
Notes: You can integrate SpikeGLX into a larger workflow with the remote scripting tools.
You can start and stop runs,
Set parameters,
Fetch and process data in real time.

## Slide 21: SpikeGLX Ecosystem (all on download page)
SpikeGLX Acquisition, Offline FileViewer
Remote SDKs Remote control / scripting
The postprocessing tools are maintained alongside SpikeGLX, they are aware of all the options and variety of data types and formats in SpikeGLX:
CatGT Filter, Remove artifacts, Extract TTL evens, Join/split runs
C_Waves Fast mean waveform calculator
TPrime ‘Aligner’ = Translate event time from stream to stream
ecephys Pipeline that runs the above and Kilosort in one package
Jennifer Colonell modified Allen Inst (Josh Siegle) version of ecephys for SpikeGLX data
Notes: Finally, SpikeGLX is an ecosystem.
Remote scripting tools
Postprocessing tools
Alignment tools